| Project Title | # Adaptive Readout Interface for Wireless Implantable Sensors |
| --- | --- |
| Student | Breandán Gillanders |
| Supervisor | Dr. Adnan Elahi |
| Co-Assessor |  |
| This Revision | Revision 1 |

# Scope Statement

Battery-less implantable LC sensors enable continuous, real-time monitoring of physiological parameters by encoding changes in variables such as pressure or strain into shifts in their resonant frequency. However, long-term performance can be affected by biological factors such as fibrotic tissue encapsulation, which alters coupling conditions, increases losses, and degrades wireless readout accuracy. Maintaining reliable operation under these changing conditions requires adaptive circuitry capable of compensating for varying link efficiency. This project aims to evaluate and improve an existing wireless readout circuit developed for passive implantable sensors. The system interrogates the implanted LC resonator via inductive coupling, measuring the reflection coefficient (S₁₁) to detect shifts in the resonant frequency of the sensor. Building on this foundation, the project will design and integrate a tuneable (reconfigurable) amplifier stage to dynamically adjust gain and stabilise performance as coupling conditions change, ensuring consistent and accurate resonance detection. The scope of the project includes the following: - Literature review on existing techniques for measuring S₁₁ and resonant frequency detection. - Circuit evaluation, schematic review, and performance benchmarking of the current system. - Simulation (LTSpice/QucsStudio) of RF front-end and amplifier stages. - Design, implementation, and tuning of a reconfigurable amplifier for adaptive gain control. - Integration and experimental validation of the improved system. - Evaluation of system stability and resonance tracking accuracy under variable coupling conditions. - Documentation and presentation of design process, results, and future recommendations.

# Stakeholder List

Stakeholder; Power; Interest; Concerns. Supervisor; High; High; Timely achievement of project milestones and solutions to project issues. Regular reporting & achievement of deliverables. Co-Assessor; Medium; Low; Achieve deliverables; clear documentation for assessment; overall project quality and completion. Student (me); High; High; Technical feasibility, meeting project scope and deadlines; demonstrating competency and producing results. Translational Medical Device Lab; High; High; Functional prototype that aligns with their research goals; practical, scalable readout system design; documentation for future development; experimental data. Technical Officers; Medium; Medium; Safe and effective use of lab equipment; availability of components. EEE Department; High; Medium; Compliance with academic standards, successful and innovative project outcome. University of Galway; High; Low; Adherence to institutional research policies and procedures.. End Users; Low; High; Reliability, safety, non-invasiveness, and ultimate improvement in patient care and monitoring.

Table of Contents

Scope Statement	1

Stakeholder List	2

Deliverables	4

Work Breakdown	4

Literature review and design requirements	4

Evaluation and refinement of the current readout circuit	5

Design of reconfigurable amplifier module	5

Reconfigurable readout circuit integration	6

Final Report	6

Work Schedule	7

Task Scheduling Table #1	7

Task Scheduling Table #2	8

Communications Strategy	9

Appendix A: Health & Safety Risk Assessments	10

Oscilloscope Risk Assessment	10

Impedance Analyser Risk Assessment	12

3D Printer Risk Assessment	14

Desk-Based Activities Risk Assessment	16

# Deliverables

Fixed Deliverables These are standard for all projects and should be incorporated into your Work Breakdown and Work Schedules below. - Project Plan (working ‘live’ document) - Progress Presentation (PPT document) - Project Presentation/Poster/Demo - Final Project Report (Word document)

Interim Deliverables - Literature Review - Existing circuit evaluation results - Circuit Simulation Results - Prototype circuits (reconfigurable amplifier circuit) - Final wireless reader system - Testing and validation results of final system

| # Work Breakdown |  |
| --- | --- |
| Work Package #1 | ## Literature review and design requirements Establish the theoretical foundation and design specifications for the resonant sensing system by reviewing relevant literature, existing technologies, and amplifier architectures. |
| Task #1 | Review the underlying theory of resonant sensing systems. |
| Task #2 | Perform a literature review on magnetic-resonant bioelectronic implants for cardiovascular monitoring, with emphasis on their readout circuitry. Additionally, review current techniques and circuit implementations for measuring reflection (S₁₁) and transmission (S₂₁) coefficients relevant to these systems. |
| Task #3 | Investigate reconfigurable and programmable amplifier designs. Define design requirements for this application (e.g. operating frequency range, max transmitted power level etc.) |
| Expected Outputs | Literature review and design requirements |

| Work Package #2 | ## Evaluation and refinement of the current readout circuit Assess and enhance the performance of the existing readout circuit through systematic testing, benchmarking, and iterative design improvements. |
| --- | --- |
| Task #1 | Review circuit schematics, PCB design and layout, and simulation models of current circuit design. Isolate and test each subsystem individually (DDS, phase shifter, phase detector, filters etc.) before performing full-system evaluation. |
| Task #2 | Perform complete system testing after individual subsystem verification. Benchmark overall performance and compare with design expectations. |
| Task #3 | Incorporate design improvements to the circuit as informed by system testing results and additional insights obtained from the literature review. |
| Expected Outputs | Existing circuit benchmark results and improved version |

| Work Package #3 | ## Design of reconfigurable amplifier module Develop, prototype, and validate a reconfigurable amplifier module tailored to the system’s design requirements for gain, bandwidth, and programmability. |
| --- | --- |
| Task #1 | Design the reconfigurable amplifier architecture based on design requirements (gain range, bandwidth etc.). Perform LTSpice simulations of the design to evaluate functionality and component selection. |
| Task #2 | Translate the simulated design into a PCB using Altium/KiCad, ensuring RF design best practices are adhered to. Fabricate the PCB in the TMD lab. Solder components to PCB. |
| Task #3 | Test the amplifier using a signal generator and an oscilloscope. Record metrics on the gain and phase response of the amplifier. Compare measured data with simulated results. |
| Expected Outputs | Validated amplifier design, PCB prototype board |

| Work Package #4 | ## Reconfigurable readout circuit integration Integrate the reconfigurable amplifier with the improved readout system, refine hardware and firmware, and produce a fully functional prototype. |
| --- | --- |
| Task #1 | Integrate the reconfigurable amplifier module with the improved readout system to evaluate system-level performance and identify requirements for merging all components onto a single PCB. Update the firmware and implement a FreeRTOS-based framework to manage concurrent tasks, including communication with the DDS, ADC, and data processing. |
| Task #2 | Develop a new PCB layout incorporating lessons learned from the existing device. Export Gerber files, BOM. Ensure components are readily available in the lab or can be easily sourced. Submit the finalized PCB design for fabrication. Procure all necessary components and assemble the revised board within the lab. |
| Task #3 | Conduct validation and performance testing of the fully integrated PCB, comparing measured results with expected results and benchmarking the final system in terms of accuracy and reproducibility. |
| Expected Outputs | Integrated PCB design and updated firmware |

| Work Package #5 | ## Final Report Compile comprehensive project documentation, analyse results, and present conclusions, limitations, and recommendations for future development. |
| --- | --- |
| Task #1 | Document the full system, detailing the circuit and PCB design, firmware implementation, and experimental results obtained at each stage of testing. |
| Task #2 | Identify limitations in the circuit design, determine areas for improvement, and compile a list of recommendations for the next design revision. |
| Task #3 | Prepare the final project report and presentation, summarising the design process, results, and key findings. Incorporate the literature review, system documentation and suggestions for future work. |
| Expected Outputs | Final system testing results |

# Work Schedule

| ## Task Scheduling Table #1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| WP # | Task # | Dependencies | Start Date | End Date | Deliverable | Notes |
| WP1 |  | n/a | 13/10/2025 | 17/11/2025 |  |  |
|  | T1 | n/a | 13/10/2025 | 22/10/2025 |  |  |
|  | T2 | n/a | 22/10/2025 | 31/10/2025 | Literature Review |  |
|  | T3 | n/a | 03/11/2025 | 17/11/2025 |  |  |
| WP2 |  |  | 22/10/2025 | 01/12/2025 |  |  |
|  | T1 |  | 22/10/2025 | 31/10/2025 |  |  |
|  | T2 |  | 03/11/2025 | 17/11/2025 | Benchmarking Results |  |
|  | T3 |  | 17/11/2025 | 01/12/2025 | Circuit Improvements |  |
| WP3 |  |  | 10/11/2025 | 08/12/2025 |  |  |
|  | T1 |  | 10/11/2025 | 24/11/2025 |  |  |
|  | T2 |  | 24/11/2025 | 01/12/2025 |  |  |
|  | T3 |  | 01/12/2025 | 08/12/2025 | Finalised Amplifer Design |  |
| WP4 |  |  | 05/01/2026 | 02/03/2026 |  |  |
|  | T1 |  | 05/01/2026 | 19/01/2026 |  |  |
|  | T2 |  | 19/01/2026 | 16/02/2026 | Wireless Reader System | Dependent on PCB manufacturing time |
|  | T3 |  | 16/02/2026 | 02/03/2026 | Validation Results |  |
| WP5 |  |  | 19/01/2026 | 30/03/2026 |  |  |
|  | T1 |  | 19/01/2026 | 16/02/2026 |  | Can be done in parallel with PCB manufacturing |
|  | T2 |  | 02/03/2026 | 09/03/2026 |  |  |
|  | T3 |  | 09/03/2026 | 30/03/2026 | Project Report |  |

| ## Task Scheduling Table #2 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| WP # | Task # | Deliv # | Sept | Oct | Nov | Dec | Jan | Feb | Mar | April |
| WP1 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  | XX |  |  |  |  |  |  |
|  | T2 |  |  | XX |  |  |  |  |  |  |
|  | T3 |  |  |  | XX |  |  |  |  |  |
| WP2 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  | XX |  |  |  |  |  |  |
|  | T2 |  |  |  | XX |  |  |  |  |  |
|  | T3 |  |  |  | XX |  |  |  |  |  |
| WP3 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  |  | XX |  |  |  |  |  |
|  | T2 |  |  |  | XX |  |  |  |  |  |
|  | T3 |  |  |  |  | XX |  |  |  |  |
| WP4 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  |  |  |  | XX |  |  |  |
|  | T2 |  |  |  |  |  | XX | XX |  |  |
|  | T3 |  |  |  |  |  |  | XX |  |  |
| WP5 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  |  |  |  | XX | XX |  |  |
|  | T2 |  |  |  |  |  |  |  | X |  |
|  | T3 |  |  |  |  |  |  |  | XXX |  |

# Communications Strategy

Name; Recipient(s); Frequency; Medium; Contents. Status Reports; Supervisor; Weekly (Friday); SharePoint; Bullet-point summary of weekly activities, issues encountered and planned next steps.. Supervisor Meetings; Supervisor; Weekly (Wednesday); In-Person; Discussion of progress, review of challenges, and identification of solutions or required support.. Progress Presentation; Supervisor & Co-Assessor; Once (start of semester 2); In-Person; Presentation summarising progress to date and outlining objectives and timelines for Semester 2.. Poster; All Stakeholders; Once (public event); Poster Session; Presentation summarising progress to date and outlining objectives and timelines for Semester 2.. Final Project Report; Supervisor & Co-Assessor; Once (end of project); PDF Document; Comprehensive final report detailing all project work, results, and conclusions..

# Appendix A: Health & Safety Risk Assessments

## Oscilloscope Risk Assessment

Person(s) at risk           Operator/instructor:  		 Students   		   Other

X

X

| Hazard(s) | Likelihood (Controls in place) | Severity (Controls in place) | Controls (See over for more details) |
| --- | --- | --- | --- |
| Electric shock | 1 | 2 | The equipment must be properly grounded. |
| Overheating of electronic components | 2 | 1 | Ensure that the component is not defective and all connections are as intended, no short circuit and all instruments are properly grounded. Use of heat sinks where needed. |

Severity        1        2        3

Likelihood:

1= Very Unlikely/Yearly

2= Unlikely/During a Semester

3= Likely/Weekly

4= Very Likely/Daily

Likelihood

4       3     2     1

Comments:

- Place an “x” in the matrix for each hazard

|  | x |  |
| --- | --- | --- |
| x |  |  |

Severity:

1= Slight Harm

2= Moderate Harm

3= Extreme Harm

| Risk Assessment with controls (Taken as the highest colour in matrix) | High | Medium | X Low |
| --- | --- | --- | --- |

Hierarchy of Controls: Equipment: Battery Powered Hand Drill Engineering: Training given for protocol, safety (PPE, spills, sharps, biohazards, waste). PPE worn at all times. Administrative/Procedural: - Standard Operating Procedures in place - Induction to lab safety provided before experiments commence - Wear a lab coat depending on your sample type. List any specific PPE required- All experiments: Lab Coat

General Rules:

- No Loose Hair, No Loose Clothes, No dangling Jewellery

- Adequate Footwear required (No open toes)

- Wash your hands after the laboratory session.

- No Eating or Drinking during laboratory sessions.

-Lab coat and gloves do not leave wet lab area.

## Impedance Analyser Risk Assessment

X

X

Person(s) at risk           Operator/instructor:  		 Students   		   Other

| Hazard(s) | Likelihood (Controls in place) | Severity (Controls in place) | Controls (See over for more details) |
| --- | --- | --- | --- |
| Liquid spills/splashes | 3 | 1 | Electrical equipment is raised & splash guard between stations. Spill tray used with large volumes of liquid. PPE worn. Spill training given, and tools provided. |
| Contact/contamination with biological tissues | 2 | 1 | PPE worn. Clean station after use. Tissue preparation on disposable biohazard mat. |
| Blood spillage/splatter | 4 | 1 | PPE worn (goggles, gloves, coat). Splash guard between stations. Clean station after use. |
| Broken glass (beakers, etc) | 2 | 1 | Spill training given, and tools provided. Training given on accident protocol. |
| Sharps injuries | 2 | 1 | Use scalpel one time only – do not clean. Dispose in sharps bin. Training given on accident protocol. |

Severity        1        2        3

Likelihood

4       3     2     1

| xxx |  |  |
| --- | --- | --- |
| x |  |  |
| x |  |  |

Likelihood:

1= Very Unlikely/Yearly

2= Unlikely/During a Semester

3= Likely/Weekly

4= Very Likely/Daily

Comments:

- Place an “x” in the matrix for each hazard

Severity:

1= Slight Harm

2= Moderate Harm

3= Extreme Harm

| Risk Assessment with controls (Taken as the highest colour in matrix) | High | Medium | X Low |
| --- | --- | --- | --- |

Hierarchy of Controls: Equipment: Prototyping Hand Tools Engineering: Training given for measurement protocol, safety (PPE, spills, sharps, biohazards, waste). PPE worn at all times. Administrative/Procedural: - Standard Operating Procedures in place - Induction to lab safety and measurement protocol provided before experiments commence - Clean Work surfaces after each lab session - Hepatitis B vaccine and titre testing offered to all workers (risks and benefits detailed) List any specific PPE required- All experiments: Gloves & Lab Coat Experiments with Blood or chemicals that require eye protection: Eye Protection

General Rules:

- No Loose Hair, No Loose Clothes, No dangling Jewellery

- Adequate Footwear required (No open toes)

- Wash your hands after the laboratory session.

- No Eating or Drinking during laboratory sessions.

-Lab coat and gloves do not leave wet lab area.

| Risk Assessment with controls (Taken as the highest colour in matrix) | High | Medium | X Low |
| --- | --- | --- | --- |

## 3D Printer Risk Assessment

Person(s) at risk           Operator/instructor:  		 Students   		   Other

| Hazard(s) | Likelihood (Controls in place) | X Severity (Controls in place) | X Controls (See over for more details) |
| --- | --- | --- | --- |
| Contact with moving parts - Entrapment or entanglement | 1 | 2 | Enclosed system guarding around moving parts Training and SOP |
| Direct/Indirect contact with parts - Electric Shock | 1 | 2 | Training and SOP Repairs and maintenance by competent person only Regular maintenance & Electrical Inspection Test & Tag |
| Exposure to fumes and ultrafine particles - Potential health effects on respiratory system from long term exposure | 2 | 1 | Place the printer away from workstations Adequate ventilation and exhaust capability Training and SOP Use of personal protective equipment (mask, gloves etc.) Regular changing of High Efficiency Particulate Arrestance HEPA filters |
| Contact with particles generated from post processing (e.g. sanding) of printed object - Potential health effects on respiratory system from long term exposure | 2 | 1 | Training and SOP Use of personal protective equipment (mask, gloves, safety glasses etc.) |
| Contact with caustic bath or other chemicals during post processing - Burns | 2 | 1 | Training and SOP Use of personal protective equipment (mask, gloves, safety glasses etc.) |
| Heat from extrusion head - Burns, scalds and other injuries by possible contact with objects and materials while they are still hot | 1 | 2 | Guarding system that prevents user from touching the nozzle or the end product until it has cooled. Training and SOP |
| Unstable surface – Printer may fall | 1 | 2 | Appropriate setup |
| Printer catching fire due to electric fault | 1 | 2 | Repairs and maintenance by competent person only Repairs and maintenance by competent person only |

Severity:

1= Slight Harm

2= Moderate Harm

3= Extreme Harm

Likelihood:

1= Very Unlikely/Yearly

2= Unlikely/During a Semester

3= Likely/Weekly

4= Very Likely/Daily

Severity        1        2        3

Likelihood

4       3     2     1

Comments:

- Place an “x” in the matrix for each hazard

|  | xxxx |  |
| --- | --- | --- |
| xxxx |  |  |

| Risk Assessment with controls (Taken as the highest colour in matrix) | High | Medium | X Low |
| --- | --- | --- | --- |

General Rules:

- No Loose Hair, No Loose Clothes, No dangling Jewellery

- Adequate Footwear required (No open toes)

- Wash your hands after the laboratory session.

- No Eating or Drinking during laboratory sessions.

-Lab coat and gloves do not leave wet lab area.

Hierarchy of Controls: Equipment: Prototyping Hand Tools Engineering: Training given for measurement protocol, safety (PPE, spills, sharps, biohazards, waste). PPE worn at all times. Administrative/Procedural: - Standard Operating Procedures in place - Induction to lab safety and measurement protocol provided before experiments commence - Clean Work surfaces after each lab session List any specific PPE required- All experiments: Gloves Lab Coat Eye Protection

| Risk Assessment with controls (Taken as the highest colour in matrix) | High | Medium | X Low |
| --- | --- | --- | --- |

## Desk-Based Activities Risk Assessment

| Hazard(s) | Likelihood (Controls in place) | Severity (Controls in place) | Controls (See over for more details) |
| --- | --- | --- | --- |
| Ergonomics of use | 1 | 1 | If using a computer habitually as a significant part of normal duty (for continuous spells of three hours or more daily) a DSE assessment must be completed. If using a laptop, a monitor should use be used with separate mouse and keyboard at your workstation. Place the laptop/notebook on a firm surface (not on the lap) at the right height for keying. Ensure that a suitable 5-star chair that is adjustable has been used to support your back. |
| Vision Fatigue | 1 | 1 | Ensure that frequent breaks and time away from the workstation for every hour of use. If glasses are required for computer work, ensure that they are been used |
| Slips, Trips and falls | 1 | 1 | Workstations should be kept clear of cables and other trip hazards |

|  |  | P |  |  |  |  | P |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| State standard of PPE: |  |  |  |  |  |  |  |  |
|  | Hard hat | Safety Glasses |  |  |  | Toe protector | Sharp protection | For all in vicinity |

| Risk Assessment Matrix |  |
| --- | --- |
| Severity. 1; 2; 3; 4. Likelihood; 1; xxxx xxxx; xxxxx xxxx xxx; x. 2; xxxx xxx. 3. 4. | Risk Assessment with controls (Taken as the highest colour in matrix). High; Medium; X Low. |