| Project Title | # Near-field Inductive Reader for Wireless, Battery-less Implantable Sensors |
| --- | --- |
| Student | Breandán Gillanders |
| Supervisor | Dr. Adnan Elahi |
| Co-Assessor |  |
| This Revision | Revision 1 |

# Scope Statement

Implantable passive LC sensors enable continuous, battery less monitoring of physiological parameters by encoding changes in variables such as pressure or strain into shifts in their resonant frequency. However, long-term performance can be affected by biological factors such as fibrotic tissue encapsulation, which alters coupling conditions, increases losses, and degrades wireless readout accuracy. Maintaining reliable operation under these changing conditions requires adaptive circuitry capable of compensating for varying link efficiency. This project aims to evaluate and improve an existing wireless readout circuit developed for passive implantable sensors. The system interrogates the implanted LC resonator via inductive coupling, measuring reflection coefficients (S₁₁) to detect resonance frequency and quality factor. Building on this foundation, the project will design and integrate a tuneable (reconfigurable) amplifier stage to dynamically adjust gain and stabilize performance as coupling conditions change, ensuring consistent and accurate resonance detection. The scope of the project includes the following: - Circuit evaluation, schematic review, and performance benchmarking of existing system. - Simulation (LTSpice/QucsStudio) of RF front-end end amplifier stages. - Design, implementation and tuning of a reconfigurable amplifier for adaptive gain control. - Integration and experimental validation of the improved system. - Evaluation of system stability and resonance tracking accuracy under variable coupling conditions. - Documentation and presentation of design process, results, and future recommendations.

# Stakeholder List

Stakeholder; Power; Interest; Concerns. Supervisor; High; High; On time & achieves milestones/outputs Timely solutions to project issues Regular Reports & achieve deliverables. Co-Assessor; Medium; Low; Achieve deliverables; clear documentation for assessment; overall project quality and completion. Student (me); High; High; Technical feasibility, meeting project scope and deadlines; demonstrating competency and producing results. Translational Medical Device Lab; High; High; Functional prototype that aligns with their research goals; practical, scalable readout system design; documentation for future development; experimental data. Technical Officers; Medium; Medium; Safe and effective use of lab equipment; availability of components. EEE Department; High; Medium; Compliance with academic standards, successful and innovative project outcome. University of Galway; High; Low; Adherence to institutional research policies and procedures.. End Users; Low; High; Reliability, safety, non-invasiveness, data privacy, and ultimate improvement in patient care or monitoring.

# Table of Contents

Scope Statement	1

Stakeholder List	1

Table of Contents	2

Deliverables	3

Work Breakdown	3

Background Research & Foundation Techniques	3

Feasibility testing/Proof-of-Concept	4

Core Project Implementation	4

Evaluations & Refinements of Core Implementation	4

Enhancements & Additional Achievements	4

Work Schedule	5

Task Scheduling Table #1	5

Task Scheduling Table #2	5

Health & Safety Risk Assessments	6

Risk Register & Risk Mitigation Plan	7

Communications Strategy	8

Optional Elements (Not Required for most Student Projects)	9

Project Budget	9

Procurement & Logistics Plan	9

Quality Assurance Plan	9

Human Resources Plan	9

Standards & Regulatory Aspects	9

Ethics & Ethical Review/Approvals	9

Data Protection and Personal Data	9

Sustainability and Environmental Impacts	9

Applicable Standards or Best Practices	9

# Deliverables

Fixed Deliverables These are standard for all projects and should be incorporated into your Work Breakdown and Work Schedules below. - Project Plan (working ‘live’ document) - Progress Presentation (PPT document) - Project Presentation/Poster/Demo - Final Project Report (Word document)

Interim Deliverables - Literature Review - Existing circuit evaluation results - Circuit Simulation Results - Prototype circuits (reconfigurable amplifier circuit) - Final wireless reader system - Testing and validation results of final system

| # Work Breakdown |  |
| --- | --- |
| Work Package #1 | ## Background Research & Foundation Techniques |
| Task #1 | Review the underlying theory of resonant sensing systems |
| Task #2 | Investigate reconfigurable and programmable amplifier designs. |
| Task #3 | Perform literature review on inductive and magnetic-resonant bioelectronic implants for cardiovascular monitoring, with emphasis on their readout circuitry. Additionally, review current techniques and circuit implementations for measuring reflection (S₁₁) and transmission (S₂₁) coefficients relevant to these systems. |
| Linked Deliverables |  |
| Expected Outputs | Literature Review |

| Work Package #2 | ## Evaluation of the Existing Device |
| --- | --- |
| Task #1 | Review circuit schematics, PCB design and layout, simulation models of existing device design |
| Task #2 | Isolate and test each subsystem individually (DDS, phase shifter, phase detector, filters etc.) before performing full-system evaluation. |
| Task #3 | Perform complete system testing after individual subsystem verification. Benchmark overall performance and compare with design expectations. |
| Linked Deliverables |  |
| Expected Outputs | Existing circuit benchmark results |

| Work Package #3 | ## Design of Reconfigurable Amplifier Module |
| --- | --- |
| Task #1 | Design the reconfigurable amplifier architecture based on performance requirements (gain range, bandwidth etc.). Perform LTSpice simulations of design to evaluate functionality and component selection. |
| Task #2 | Translate the simulated design into a PCB using Altium/KiCad, ensuring RF design best practices are adhered to. Fabricate PCB in TMD lab. Solder components to PCB. |
| Task #3 | Test the amplifier using signal generator and oscilloscope. Record metrics on gain and phase response of the amplifier. Compare measured data with simulated results. |
| Linked Deliverables |  |
| Expected Outputs | Validated amplifier design, PCB prototype board |

| Work Package #4 | ## Revised Device Design |
| --- | --- |
| Task #1 | Integrate the reconfigurable amplifier and any additional circuitry needed (additional filtering etc.) into the system schematics. |
| Task #2 | Develop a new PCB layout incorporating lessons learned from the existing device. Export Gerber files, BOM. Ensure components are readily available in the lab or can be easily sourced. |
| Task #3 | Submit the finalized PCB design for fabrication. Procure all necessary components and assemble the revised board within the lab. |
| Task #4 | Update the software to ensure compatibility with revised hardware. Implement a FreeRTOS based framework to manage concurrent tasks such as communication with DDS, ADC, and data processing. |
| Linked Deliverables |  |
| Expected Outputs | Revised design PCB |

| Work Package #5 | ## Validation of Revised Design |
| --- | --- |
| Task #1 | Conduct initial continuity and power checks on the new board. |
| Task #2 | Calibrate the revised system under controlled conditions. Test different reader coils at various lengths to the sensor to determine the optimal configuration. |
| Task #3 | Evaluate the complete system by varying the sensor output under controlled conditions (e.g., placing the sensor in a pressure chamber and comparing readings against known reference pressures) Assess the accuracy and reproducibility of the results. |
| Linked Deliverables |  |
| Expected Outputs | Final system testing results |

# Work Schedule

| ## Task Scheduling Table #1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| WP # | Task # | Dependencies | Start Date | End Date | Deliverable | Notes |
| WP1 |  | n/a | 06/10/2025 | 07/11/2025 |  |  |
|  | T1 | n/a | 06/10/2025 | 20/10/2025 |  | Completed |
|  | T2 | n/a | 20/10/2025 | 07/11/2025 |  |  |
|  | T3 | n/a | 22/10/2025 | 07/11/2025 |  |  |
| WP2 |  | WP1 | 13/10/2025 | 10/11/2025 |  |  |
|  | T1 | WP1 T2 | 13/10/2025 | 20/10/2025 |  | Completed |
|  | T2 | WP2 T1 | 20/10/2025 | 27/10/2025 |  |  |
|  | T3 | WP2 T2 | 27/10/2025 | 10/11/2025 |  |  |
| WP3 |  | n/a | 10/11/2025 | 08/12/2025 |  |  |
|  | T1 | WP1 T3 | 10/11/2025 | 17/11/2025 |  |  |
|  | T2 | WP3 T1 | 17/11/2025 | 01/12/2025 |  |  |
|  | T3 | WP3 T2 | 01/12/2025 | 08/12/2025 |  |  |
| WP4 |  | WP3 | 05/01/2026 | 16/02/2026 |  |  |
|  | T1 | WP3 | 05/01/2026 | 12/01/2026 |  |  |
|  | T2 | WP4 T1 | 12/01/2026 | 26/01/2026 |  |  |
|  | T3 | WP4 T2 | 26/01/2026 | 16/02/2026 |  | Dependent on PCB manufacturing time |
|  | T4 | WP3 | 26/01/2026 | 16/02/2026 |  | In parallel with PCB manufacturing |
| WP5 |  | WP4 | 16/02/2026 | 02/03/2026 |  |  |
|  | T1 | WP4 | 16/02/2026 | 18/02/2026 |  |  |
|  | T2 | WP5 T1 | 18/02/2026 | 23/02/2026 |  |  |
|  | T3 | WP5 T2 | 23/02/2026 | 02/03/2026 |  |  |

| ## Task Scheduling Table #2 |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| WP # | Task # | Deliv # | Sept | Oct | Nov | Dec | Jan | Feb | Mar | April |
| WP1 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  | XX |  |  |  |  |  |  |
|  | T2 |  |  | XX |  |  |  |  |  |  |
|  | T3 |  |  | XX | X |  |  |  |  |  |
| WP2 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  | X |  |  |  |  |  |  |
|  | T2 |  |  | X |  |  |  |  |  |  |
|  | T3 |  |  | X | X |  |  |  |  |  |
| WP3 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  |  | XX |  |  |  |  |  |
|  | T2 |  |  |  | XX |  |  |  |  |  |
|  | T3 |  |  |  |  | XX |  |  |  |  |
| WP4 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  |  |  |  | XX |  |  |  |
|  | T2 |  |  |  |  |  | XX |  |  |  |
|  | T3 |  |  |  |  |  | X | XX |  |  |
|  | T4 |  |  |  |  |  |  | XX |  |  |
| WP5 |  |  |  |  |  |  |  |  |  |  |
|  | T1 |  |  |  |  |  |  | X |  |  |
|  | T2 |  |  |  |  |  |  | X |  |  |
|  | T3 |  |  |  |  |  |  | X | X |  |

# Health & Safety Risk Assessments

[Note – standard activity risk assessments attached as Appendix A]

# Risk Register & Risk Mitigation Plan

# Communications Strategy

Name; Recipient(s); Frequency; Medium; Contents. Status Reports; Supervisor; Weekly (Friday); SharePoint; Bullet point summary of activities, blocking issues, etc.. Supervisor Meetings; Supervisor; Weekly (Wednesday); In Person; Summary of progress during the period; note any blocking issues & solution or still open. Progress Presentation; Supervisor & Co-Assessor; Once (deliverable); In Person; Presentation on progress made on project in semester 1 and plans to complete the project for semester 2.. Poster; All Stakeholders; Once (deliverable); Public Event; Visual summary of project objectives, results etc.. Project Report; Supervisor & Co-Assessor; Once (deliverable); PDF Document; Final project report.